## Instructions
Title: TD,NovaSeq,MATL-RPLC-QUAL
ER Number: 1034689
1 | Purpose
The purpose of this Training Document is to provide qualified personnel with required guidelines to verify material replacement or reinstallation efficacy and quality on a NovaSeq instrument.
2 | Scope
This training document is applicable for NovaSeq 6000 products that are qualified per reference training documents cited in Section 3. This document is applicable for qualifying assemblies/sub-assemblies in the Qualification Table . Material not cited (newly qualified, non-released, under deviation, spares and refurbs), but are of an equivalent category are still applicable to this document. This document does not provide instructions on how to troubleshoot or replace material. This training document is comprised of two sections on separate tabs within this excel spreadsheet: 1) Instructions, and 2) Qualification Table. The Instructions section provides references and document usage, while the Table provides a matrix of commonly replaced or reinstalled parts and required operational steps to be performed. Users of this document need to be trained and adhere to all corresponding NovaSeq Manufacturing Operation Training Guides and relevant Work Instructions. Manufacturing and Life Cycle Management is responsible for maintaining and updating this Training Document.
3 | References
| DIR |  | Document Description
| 1000000061155 |  | PROC,NovaSeq,iMFG-PROCESS
| 1000000028564 |  | TD,NovaSeq,WITS-GUIDE
| 1000000020726 |  | TD,NovaSeq,FIT,MOCK-RUN
| 1000000061160 |  | TD,NovaSeq,INTEGRATION,BIM-RCA-CNNCT
| 1000000061161 |  | TD,NovaSeq,INTEGRATION,XYA-DAD-CTA-IOM
| 1000000061162 |  | TD,NovaSeq,INTEGRATION,OAS-ENC-BIRD
| 1000000061164 |  | TD,NovaSeq,FUNCTIONAL,INIT-POWER
| 1000000024684 |  | TD,NovaSeq,JOIN-DOMAIN
| 1000000061166 |  | TD,NovaSeq,FLUIDICS
| 1000000020724 |  | TD,NovaSeq,ALIGNMENT
| 1000000020482 |  | TD,NovaSeq,FC-THERM_CAL
| 1000000020491 |  | TD,NovaSeq6k,FIT
| 1000000040613 |  | TD,NovaSeq,FIT,FLO
| 1000000073101 |  | SPEC,NovaSeq6k,FIT,FIT_SPECIFICATIONS
4 | Instructions
| 1) | The Table section provides applicable materials and assemblies within the top header. Each subsequent row provides the operation (FUNC OP thru FIT OP) and associated steps. These steps are mirrored from each operational training document's table of contents. Steps identified with an asterix (*) are not listed within the MES rework workflow but are needed for proper qualification. Note, each step within the training document is a high level description in which not all coinciding steps may be applicable for qualifying the material. Determination of applicability must be reviewed by one or more of the following personnel: Manufacturing Lead, Process Owner, Supervisor, or Manager. Removal or addition of steps must be documented and justified appropriately.
| 2) | Users of this document are to first identify the applicable material or assembly and follow the recommended operational steps marked. See legend below for marking definitions.
Not Required
|  |  |  | Required |  | X
|  |  |  | Optional (dependent on instrument results) |  | ?
Required operational steps only need to be followed up until the original point of observed failure and material replacement or reinstallation. For example, a material replaced or reinstalled within the fluidics operations should only be qualified to the equivalent fluidics operations. Pending all qualifying testing passes the standard acceptance critieria, the instrument should then continue with normal manufacturing in-process testing. For every training document required to qualify the material and instrument, any provided ESD, Safety, and General Precautions must be followed prior to working on the instrument. Each training document cited within this document may coincide to a particular work instruction document based on product. The completion of any applicable work instructions during the replaced material qualification is required, and any deviations is at the discretion of Manufacturing Lead, Process Owner, Supervisor, or Manager that is approving the process and release of the instrument. Removal or absence of any work instructions required to qualify the material replacement must be documented and justified appropriately.
| 3) | During instances in which qualification of the replaced or reinstalled material fails, the normal troubleshooting and evaluative processes should be used to assess failure. Qualification of the subsequently replaced or reinstalled material can be assessed via this document as needed. During instances when the replaced or reinstalled material cannot be completely qualified per recommended guidelines due to a dependent failure in a different material requiring replacement or reinstallation, the qualification should still continue in parallel with the subsequently replaced or reinstalled material to complete qualification of all materials. Original problem statement must be verified before operator moves to next task.
5 | Release History
Version |  | ER # | Originator |  | Description of Change
00 |  | 1032317 | D.Gates |  | Initial Release. Sheet protection required, with exception to "Use Autofilter" for Table section.
01 |  | 1034689 | V. Tan |  | In Instructions tab: -Reference section, removed DHR Coversheets, Uncrating, MAT docs, and FII documents -Added verbiage to include reinstallation of material and updated verbiage to align with current practice. In Qualification tab: -Updated formatting for Materials for better ease of view -Removed CA MAT qualification steps due to process update and added to FUNC qualification steps. -Added MES Rework Steps column. -Updated steps to align with current practice. -Added BIM Drawer, Common line, and Light Band to materials. -Added necessary qualification steps for Degasser. -Added note for Flowrate Sensor to only perform test for SOUR2 position. -Added note for Degasser to only perform test for VSM and VB2. -Added note for BIM Drawer replacement for BIM test to only check drawer sensors. -Removed BIRD Ubertarget Alignment qualification step for Camera and LGM.

## Qualification Table
Instrument Qualification Table
OP | MES Rework Steps | Training Document Steps | MIB BOARD | FCH BOARD | FIB BOARD | SBC BOARD | SYS BOARD | SYRINGE PUMP | FLOW RATE SENSOR | FLUIDICS MODULE (FLM) | REAGANT CHILLER ASSEMBLY (RCA) | BUFFER INTERFACE MODULE (BIM) | BIM DRAWER | DEGASSER | ASSY, TUBING KIT, COMMON LINES, SLEEVE | DUAL ACTUATION DECK (DAD) | CABLE TRACK ASSEMBLY (CTA) | ASSY, XY STAGE MODULE (XYA) | TIP TILT ASSEMBLY (TTA) | ASSY, CAMERA MODULE (CAM) | ASSY, FOCUS TRACKING MODULE (FTM) | ASSY, EMISSION OPTICS MODULE (EOM) | LGM_HT_V2 ACTUATORS (LGM) | BIRD UBERTARGERT | Z-STAGE MOTION CONTROLLER | FC ENCLOSURE | ASSY, LIGHTBAND | ASSY, OPA W/ NOZZLE (OPA) | COMPUTE ENGINE (CE) | OPTICAL AIR SHIELD (OAS) | COOLANT RESERVOIR TRAY ASSEMBLY | PSU
FUNC | System Setup | Verify SBC Image |  |  |  | X
FUNC | System Setup | Join Domain |  |  |  | X
FUNC | System Setup | Copy current software package into designated local folder |  |  |  | X
FUNC | System Setup | Reference file in software package for upgrading software |  |  |  | X
FUNC | System Setup | Use NovaSeqXMLSpecFileUtility Tool |  |  |  | X
FUNC | System Setup | Set Output Repository before unsimulation |  |  |  | X
FUNC | System Setup | Update Instrument Serial Number in Software |  |  |  | X
FUNC | System Setup | Verify Instrument Type |  |  |  | X
FUNC | System Setup | Update Override Configuration |  |  |  | X
FUNC | System Setup | Update FW and power cycle instrument |  | x | x | x | x |  |  |  | x | x | [16 empty] | x |  |  | x
FUNC | System Setup | Verify NVSS Initializes without unexpected errors or warnings* | x | x | x | x | x | x | x | x | x | x |  | x | x | x | x | x | x | x | x | x | x |  | x | x | x | x | x | x | x | x
FUNC | System Setup | Set Output Repository |  |  |  | x
FUNC | System Setup | Verify NVSS version |  |  |  | x
FUNC | System Setup | Verify CloudAPIKey | [29 empty] | x
FUNC | System Setup | Verify CE LED lights are on |  |  |  | x | [22 empty] | x
FUNC | System Setup | Verify CE Ethernet cable connection is green |  |  |  | x | [22 empty] | x
FUNC | System Setup | Verify Z: drive is active and D: is present |  |  |  | x | [22 empty] | x
FUNC | System Setup | Verify no coolant leaks are present around optical modules | [17 empty] | x | x | x | x
FUNC | Light Band Burn-In | Perform Light Band Burn In |  | X | [22 empty] | x
FUNC | BIM Tests | Perform BIM Mechanical and RFID tests | [9 empty] | x | x
FUNC | RCA Tests | Perform RCA Mechanical and RFID tests | [8 empty] | x
FUNC | DAD Tests - Vacuum Pressure Sensor | Enter Clamp calibration values |  | x | [11 empty] | x
FUNC | DAD Tests - Vacuum Pressure Sensor | Perform Vacuum performance test | x | x | [11 empty] | x | x | x
FUNC | DAD Tests - Vacuum Present Sensor | Load DRY FCs and verify FC present sensor indicates present | x | x | [11 empty] | x | x | x
FUNC | DAD Tests - Vacuum Present Sensor | Verify FC vacuum ≤ -90 kPa | x | x | [10 empty] | x | x | x | X
FUNC | DAD Tests - Vacuum Present Sensor | Ensure vauum is <1.0 kPa when FCs are disengaged | x | x | [10 empty] | x | x | x | X
FUNC | DAD Tests - Vacuum Present Sensor | Verify FC present sensor indicate not present | x | x | [11 empty] | x | x | X
FUNC | ICA Burn In | Perform ICA Burn In | x | x | [6 empty] | x | x |  |  |  | x | x | [8 empty] | x | x
FUNC | XY Stage Tests | Perform XY Stage tests | x | x | x | [9 empty] | x | x | X | X | [10 empty] | x
FUNC | XY Stage Tests | Perform Tip Tilt Motor tests | X | x | [11 empty] | X | X | X | X
FUNC | XY Stage Tests | XY Motor Servo Test | x | x | x | [10 empty] | x | x | X
FUNC | Interlock Verification | Verify Interlocks |  |  | X | [20 empty] | X
FUNC | Coolant Level Sensor Check | Ensure ALL coolant levels are topped off |  |  |  |  | X |  |  |  | X |  |  |  |  | X | X |  |  | X | X |  | X |  |  |  |  | X |  |  | X
FUNC | Coolant Level Sensor Check | Ensure reservoir level sensors read “true” |  |  |  |  | X | [23 empty] | X
FUNC | System Automated Test | Perform system Thermal Automated Tests |  | X | X |  | X |  |  |  | X |  |  |  |  | x | x | [10 empty] | X |  |  | x
FUNC | System Thermal Test | Perform RCA Coolant loop test |  |  |  |  | X |  |  |  | x | [19 empty] | x
FUNC | System Thermal Test | Perform Optical Coolant Loop test |  |  | X |  | X | [12 empty] | X | X | X | X | [7 empty] | x
FUNC | FCH Thermal Test -SideA FCH Thermal Test -SideB | Perform FCH Thermal Performance | x | x | [11 empty] | x | x | [10 empty] | X
FUNC | Optics General Tab Tests | Verify Laser controller lights are on |  |  | X | [12 empty] | X |  |  |  |  | X |  |  |  |  | X
FUNC | Optics General Tab Tests | Verify EOM Internal Z-stage lock |  |  | X | [16 empty] | X |  |  | x |  |  | X
FUNC | Optics General Tab Tests | Verify good Camera images |  |  | X | [12 empty] | X |  | X |  | x | [5 empty] | X | X
FUNC | Optics General Tab Tests | Verify noise is visible on FTM camera |  |  | X | [12 empty] | X |  |  | X | x | [5 empty] | X
FUNC | Optics General Tab Tests | Verify Zoom Lens Movement |  |  | X | [12 empty] | X |  |  |  | x | [5 empty] | X
FUNC | Optics Laser Tab Tests | Verify Laser readout power can increase to 50 mW; then disable lasers |  |  | X | [12 empty] | X |  |  |  |  | X |  |  |  |  | X
FUNC | Objective Air Shield Tests | Perform OAS Test |  | X | X | [12 empty] | X |  |  |  | X | X |  |  |  |  | X |  | X
FUNC | Camera Bit Error Tests | Perform FTM Bit Error test |  |  | X | [12 empty] | X |  |  | X | X | [5 empty] | X
FUNC | Camera Bit Error Tests | Perform EOM Bit Error tests |  |  | X | [12 empty] | X |  | X |  | X | [5 empty] | X | x
FUNC | Z-Stage Tests | Perform Z-Stage tests |  |  | X | [12 empty] | X |  |  |  | X |  |  | X |  |  | X
FUNC | CE Image Verification | Ensure LCM has verified CE Image | [26 empty] | X
FUNC | CE Image Verification | Lock Internal Z-stage Lockout (Only required if exiting NVSS)* | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x
FLUIDICS | Degasser Test | Ensure tubing are connected correctly and are secured* | [5 empty] | x | x | x | x | x |  | x
FLUIDICS | Degasser Test | Perform Degasser Test | [7 empty] | X |  |  |  | X
FLUIDICS | Pressure Sensor Test | Perform Fluidics Sensor test | [7 empty] | X
FLUIDICS | Flow Rate Sensor Test | Perform Flow Rate Sensors Test | [7 empty] | X
FLUIDICS | Valve Response Test | Perform Valve response test | [8 empty] | X
FLUIDICS | Bypass Pressure Decay test | Perform bypass Pressure Decay (without flowcells) |  | X |  |  |  | X |  | X | X | X
FLUIDICS | 4-lane Pressure Decay test | Perform Flowcell Pressure Decay Test |  | X |  |  |  | X |  | X | X | X |  |  | x | X | X | X
FLUIDICS | Priming | Execute Priming from Dry prime recipe |  | X |  |  |  | X | X | X | X | X |  | X | x | X | X | X
FLUIDICS | Plumbing Verification Tests | Perform Plumbing verification Tests | [5 empty] | X | X | X | X | X |  | X |  | x | x
FLUIDICS | Volumetrics Test | Perform Volumetrics test |  | X |  |  |  | X | X | X | X | X |  | X | x | X | X | X
FLUIDICS | Impedance Test | Perform Impedance test |  | X |  |  |  | X | X | X | X | X |  | X | x | X | X | X
FLUIDICS | Impedance Test | Lock Internal Z-stage Lockout (Only required if exiting NVSS)* | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x
ALIGN | LGM Power Measurement/Set-up | Ensure zoom lens values are cleared* | [19 empty] | x
ALIGN | LGM Power Measurement/Set-up | LaserPowerReading.log and *LaserPowerResult.xml copied to Repository (LGM Power Measurement) | [15 empty] | X |  |  |  |  | X |  |  |  |  | X
ALIGN | Initial FTM Adjustment | Verify FCA and FCB Vacuum Pressure (UT Loading) | [12 empty] | X | X | X | X | X | X | X | X | X |  |  |  |  | X
ALIGN | Initial FTM Adjustment | Collided FTM spots are centered at X = 1020 and Y = 540 (Initial FTM Alignment) | [13 empty] | X | X | X | X |  | X | X | [5 empty] | X
ALIGN | Initial FTM Adjustment | FTM adjustment cover turned back, Center Screw Locked | [13 empty] | X | X | X | X |  | X | X | [5 empty] | X
ALIGN | Initial FTM Adjustment | Side A Tilt Motor Start Positions Updated and Saved in Config Editor | [13 empty] | X | X | X | X |  | X | X | [5 empty] | X
ALIGN | Initial FTM Adjustment | Coarse XY offset completed | [13 empty] | X | X | X | X |  | X | X | [5 empty] | X
ALIGN | Initial FTM Adjustment | PWM Laser Power Calibration completed | [13 empty] | X | X | X | X |  | X | X | [5 empty] | X
ALIGN | Laser Line/Fine XY Offset | Initial Laser Line Centering | [13 empty] | X | X | X | X | X |  | X | X |  |  |  |  | X
ALIGN | Laser Line/Fine XY Offset | Initial Laser Line Rotation | [13 empty] | X | X | X | X | X |  | X | X |  |  |  |  | X
ALIGN | Laser Line/Fine XY Offset | Laser Line Uniformity | [13 empty] | X | X | X | X | X |  | X | X |  |  |  |  | X
ALIGN | Laser Line/Fine XY Offset | FineXY Offset completed | [12 empty] | X | X | X | X | X | X | X | X | X |  |  |  |  | X
ALIGN | FTM Final Adjustment | FTM ROI Calibration and Focus Model Generation | [13 empty] | X | X | X | X | X | X | X | X |  |  |  |  | X
ALIGN | FTM Final Adjustment | FTM Focus Model Repeatabilty and Spot Separation Stability |  |  | X | [10 empty] | X | X | X | X | X | X | X | X |  | X |  |  | X
ALIGN | FTM Final Adjustment | FTM Boresight Error | [13 empty] | X | X | X | X | X | X | X | X |  |  |  |  | X
ALIGN | XYA Rotation and Stage Skew | X-Tilt Calibration (Ubertarget) completed | [12 empty] | X | X | X | X | X | X | X | X | X |  |  |  |  | X
ALIGN | XYA Rotation and Stage Skew | Distortion, DAD Skew and XYA Rotation | [12 empty] | X | X | X | X | X | X | X | X | X |  |  |  |  | x
ALIGN | XYA Rotation and Stage Skew | Side B Calibration and Skew Rotation | [12 empty] | X | X | X | X | X | X | X | X | X |  |  |  |  | x
ALIGN | Chromatism | UT Chromatism | [12 empty] | X | X | X | X | X | X | X | X | X |  |  |  |  | X
ALIGN | Chromatism | RFID test completed | x | X | [10 empty] | X | X | X | X | X | X | X | X | X |  |  |  |  | X
ALIGN | BIRD Alignment | BIRD Ubertarget Alignment (Tilt Motor Start Positions Updated and Saved in Config Editor + Autocenter) | [12 empty] | X | X | X | X | X |  | X | X |  | X |  |  |  | X
ALIGN | BIRD Alignment | FTM ROI Calibration and Focus Model Generation (BIRD) | [13 empty] | X | X | X | X | X | X | X | X | X |  |  |  | X
ALIGN | BIRD Alignment | X-Tilt Calibration (UBIRD) completed | [13 empty] | X | X | X | X | X | X | X | X | X |  |  |  | X
ALIGN | BIRD Alignment | BIRD Chromatism | [12 empty] | X | X | X | X | X | X | X | X | X | X |  |  |  | X
ALIGN | BIRD Alignment | EOM Stack Collection - Distortion | [12 empty] | X | X | X | X | X | X | X | X | X | X |  |  |  | X
ALIGN | Final Alignment | Verify Laser Line | [12 empty] | X | X | X | X | X | X | X | X | X | X |  |  |  | X
ALIGN | Final Alignment | Ensure UT Chromatism passes | [12 empty] | X | X | X | X | X | X | X | X | X | X |  |  |  | X
ALIGN | Final Alignment | Lock Internal Z-stage Lockout (Only required if exiting NVSS) | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x
THERM | Calibrate and Verify Side A Calibrate and Verify Side B | Load HCT and ensure vacuum engages | X | X | [11 empty] | X | X | X
THERM | Calibrate and Verify Side A Calibrate and Verify Side B | Perform Calibration and Verification | X | X | [11 empty] | X | X | X
THERM | Calibrate and Verify Side A Calibrate and Verify Side B | Lock Internal Z-stage Lockout (Only required if exiting NVSS)* | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x
FIT | Pre-FIT Activities | BIRD Measurement | [13 empty] | X | X | X | X | X | X | X | X | X
FIT | FLO Calibration | FLO | [8 empty] | X |  |  |  | x | X | X | X |  |  | X |  | X |  |  |  |  | X
FIT | FLO Calibration | Lock Internal Z-stage Lockout (Only required if exiting NVSS)* | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x | x
FIT | FIT Set up and Begin Sequencing Run FIT Run Completion Post-FIT Instrument Wash BIRD Analysis Wrap Up | FIT (Final Instrument Test) | X | X | X | X | X | X | X | X | X | X |  | X | X | X | X | X | X | X | X | X | X | X | X | X |  | X | X | X | X | X
Note on C13: Tan, Veronica: Step is not part of the MES rework workflow for eDHR records, but is necessary to perform in MFG for certain component replacements
Note on N22: Tan, Veronica: Only need to check Drawer sensors by opening and closing drawer; check status is correct (No to Yes)
Note on C52: Tan, Veronica: Step is not part of the MES rework workflow for eDHR records, but is necessary to perform in MFG if exiting NVSS.
Note on C53: Tan, Veronica: Step is not part of the MES rework workflow for eDHR records, but is necessary to perform in MFG for certain component replacements
Note on J61: Tan, Veronica: Only perform SOUR2 position
Note on O61: Tan, Veronica: Only perform VSM and VB2 positions
Note on C64: Tan, Veronica: Step is not part of the MES rework workflow for eDHR records, but is necessary to perform in MFG if exiting NVSS.
Note on C92: Tan, Veronica: Only required if exiting NVSS
Note on C95: Tan, Veronica: Step is not part of the MES rework workflow for eDHR records, but is necessary to perform in MFG if exiting NVSS.
Note on C98: Tan, Veronica: Step is not part of the MES rework workflow for eDHR records, but is necessary to perform in MFG when exiting NVSS.

## ICALelve (hidden)
ICA Level Re-testing |  |  | Integrated Chassis Assembly (ICA)
OP | Training Document Steps | EXAMPLE | DUAL ACTUATION DECK (DAD) | CABLE TRACK ASSEMBLY (CTA) | FLUIDICS MODULE (FLM) | REAGANT CHILLER ASSEMBLY (RCA) | BUFFER INTERFACE MODULE (BIM) | SLAB DOOR MECHANISM | MIB BOARD | FCH BOARD | FIB BOARD
ICA MAT | Perform Flowcell Door test | [6 empty] | X |  | X
ICA MAT | Perform Light Band test for side A and side B | [8 empty] | X
ICA MAT | Perform System Thermal Automated Tests |  | X | X |  |  |  |  | X | X
ICA MAT | Perform Flow Cell Holder Vacuum Test for side A and side B: |  | X | X |  |  |  |  | X | X
ICA MAT | Perform BIM Verification Test for side A and side B: | [5 empty] | X
ICA MAT | Perform RCA Verification Tests for side A and side B: |  |  |  |  | X
ICA MAT | Record RCA Thermal RTDs in datasheet |  |  |  |  | X
ICA MAT | Perform FCH Thermal Performance for side A and side B |  | X | X |  |  |  |  | X | X
ICA MAT | Record Fluidics Pressure Sensors in datasheet |  |  |  |  | X | X
ICA MAT | Perform Valve Response Test for side A and side B |  |  |  |  | X | X |  |  |  |  | `
ICA MAT | Perform Bypass Automated Pressure Decay (without flowcells) for side A and side B |  | X | X |  | X | X |  | X | X
ICA MAT | Perform Flow Cell Automated Pressure Decay (with flowcells) for side A and side B |  | X | X |  | X | X |  |  | X
ICA MAT | Perform Objective Air Shield (OAS) Test | [8 empty] | X
ICA MAT | Ensure all coolant reservoirs are topped off and not leaking and coolant line connections are not leaking |  | X | X

## Functional Level (hidden)
Integrated Chassis Assembly (ICA)
OP | Training Document Steps | EXAMPLE | ASSY, COMPUTE MODULE (CE) | ASSY, OPA W/ NOZZLE (OPA) | ASSY, XY STAGE MODULE (XYA) | TIP TILT ASSEMBLY (TTA) | ASSY, CAMERA MODULE (CAM) | ASSY, FOCUS TRACKING MODULE (FTM) | ASSY, EMISSION OPTICS MODULE (EOM) | LGM_HT_V2 ACTUATORS (LGM) | BIRD UBERTARGERT | DUAL ACTUATION DECK (DAD) | CABLE TRACK ASSEMBLY (CTA) | FLUIDICS MODULE (FLM) | REAGANT CHILLER ASSEMBLY (RCA) | BUFFER INTERFACE MODULE (BIM) | SLAB DOOR MECHANISM | MIB BOARD | FCH BOARD | FIB BOARD | Z-STAGE CONTROLLER BOARD
ICA MAT | Perform Flowcell Door test |  |  |  | X | [11 empty] | X
ICA MAT | Perform Light Band test for side A and side B |  |  |  | X
ICA MAT | Perform System Thermal Automated Tests |  |  |  | X | [6 empty] | X | X
ICA MAT | Perform Flow Cell Holder Vacuum Test for side A and side B: |  |  |  | X | [6 empty] | X | X
ICA MAT | Perform BIM Verification Test for side A and side B: | [14 empty] | X
ICA MAT | Perform RCA Verification Tests for side A and side B: | [13 empty] | X
ICA MAT | Record RCA Thermal RTDs in datasheet | [13 empty] | X
ICA MAT | Perform FCH Thermal Performance for side A and side B |  |  |  | X | [6 empty] | X | X
ICA MAT | Record Fluidics Pressure Sensors in datasheet
ICA MAT | Perform Valve Response Test for side A and side B | [13 empty] | X
ICA MAT | Perform Bypass Automated Pressure Decay (without flowcells) for side A and side B |  |  |  | X | [6 empty] | X | X |  | X
ICA MAT | Perform Flow Cell Automated Pressure Decay (with flowcells) for side A and side B |  |  |  | X | [6 empty] | X | X
ICA MAT | Perform Objective Air Shield (OAS) Test | [17 empty] | X
ICA MAT | Ensure all coolant reservoirs are topped off and not leaking and coolant line connections are not leaking |  |  |  | X | [6 empty] | X | X
FUNC | Change Instrument SN on Computer and Doman
FUNC | Copy current software package into C:\Illumina folder
FUNC | Reference file in software package for upgrading software
FUNC | Verify NVSS Initializes without unexpected errors or warnings | X
FUNC | Set Output Repository
FUNC | Update Instrument Serial Number in Software
FUNC | Verify Instrument Type
FUNC | Update Override Configuration
FUNC | Un-simulate components previously simulated
FUNC | Update FW and power cycle instrument
FUNC | Set Output Repository
FUNC | Verify NVSS Initializes without unexpected errors or warnings
FUNC | Verify NVSS version | X
FUNC | Verify Firmware version
FUNC | Verify no warnings on Mock Run page and notify LCM to verify CE image
FUNC | Verify CE LED lights are on
FUNC | Verify CE Ethernet cable connection is green | X
FUNC | Verify Z: drive is active and D: is present
FUNC | Verify no coolant leaks are present around optical modules
FUNC | Generate CloudAPIKey
FUNC | Perform XY Stage tests |  |  | X | X | X | [6 empty] | X
FUNC | Perform Tip Tilt Motor tests |  |  | X | X | X | [6 empty] | X
FUNC | XY Motor Servo Test (20 cycles) |  |  | X | X | X | [6 empty] | X
FUNC | Ensure FCs and FCHs are clean
FUNC | Load/Unload DRY FCs
FUNC | Ensure to set Descriptors to NovaSeq S3
FUNC | Verify FC vacuum < -90 kPa |  |  | X | X | [6 empty] | X | X
FUNC | Perform Pressure Decay Test | X |  | X | X | [6 empty] | X | X
FUNC | Ensure Flowcell Holder door and Optical Module cover are closed | X |  | X
FUNC | Verify (4) LEDs on FIB are enabled |  |  | X
FUNC | Verify Laser interlocks are enabled |  |  | X
FUNC | Ensure ALL coolant levels are topped off |  |  | X
FUNC | Ensure reservoir level sensors read “true” |  |  | X
FUNC | Relaunch software |  |  | X
FUNC | Perform Optical Coolant Loop test |  |  | X
FUNC | Verify Laser controller lights are on |  |  | X | [15 empty] | X
FUNC | Verify good Camera images |  |  | X | [15 empty] | X
FUNC | Verify noise is visible on FTM camera |  |  | X
FUNC | Verify Zoom Lens Movement |  |  | X | [15 empty] | X
FUNC | Verify Laser readout power can increase to 50 Mw |  |  | X | [15 empty] | X
FUNC | Disable lasers |  |  | X | [15 empty] | X
FUNC | Verify Laser Shutter functionality |  |  | X | [15 empty] | X
FUNC | Perform OAS Test; Verify passing results |  |  | X | [15 empty] | X
FUNC | Perform FTM Bit Error test |  |  | X | [14 empty] | X
FUNC | Perform EOM Bit Error tests |  |  | X | [15 empty] | X
FUNC | Perform Z-Stage tests |  |  | X | [15 empty] | X
FUNC | Ensure LCM has verified CE Image | [18 empty] | X
FLUIDICS | Launch NVSS and verify that NVSS initializes without errors | X | [9 empty] | X | X
FLUIDICS | Execute Priming from Dry prime recipe for Side A and Side B | [10 empty] | X | X |  |  |  |  | X | X
FLUIDICS | Perform Volumetrics test for Side A and Side B | X | [9 empty] | X | X |  |  |  |  | X | X
FLUIDICS | Ensure Volumetrics Test Results Pass specification | X | [9 empty] | X | X |  |  |  |  | X | X
FLUIDICS | Perform Impedance test for Side A and Side B | X | [9 empty] | X | X |  |  |  |  | X | X
FLUIDICS | Ensure Impedance Test Results Pass specification | X | [9 empty] | X | X |  |  |  |  | X | X
Note on B42: Liao, Uland: Revert temperature calibrations to default values in config file. Date: 6/30/2016

## Optics_Therm_FIT_Table (hidden)
Integrated Chassis Assembly (ICA)
OP | Training Document Steps | EXAMPLE | ASSY, COMPUTE MODULE (CE) | ASSY, OPA W/ NOZZLE (OPA) | ASSY, XY STAGE MODULE (XYA) | TIP TILT ASSEMBLY (TTA) | ASSY, CAMERA MODULE (CAM) | ASSY, FOCUS TRACKING MODULE (FTM) | ASSY, EMISSION OPTICS MODULE (EOM) | LGM_HT_V2 ACTUATORS (LGM) | Z-STAGE CONTROLLER BOARD | DUAL ACTUATION DECK (DAD) | CABLE TRACK ASSEMBLY (CTA) | FLUIDICS MODULE (FLM) | REAGANT CHILLER ASSEMBLY (RCA) | BUFFER INTERFACE MODULE (BIM) | SLAB DOOR MECHANISM | MIB BOARD | FCH BOARD | FIB BOARD
ALIGN | NVSS Initializes without Errors | X |  | X | X |  | X | X | X | X | X | X | X
ALIGN | OPA Brace/OAS Nozzle reinstalled and (8) screws torqued to 50 in-lbs |  |  | X | X |  |  |  |  | X
ALIGN | LaserPowerReading.log and *LaserPowerResult.xml copied to Repository |  |  | X | X |  |  |  |  | X
ALIGN | Verify FCA and FCB Vacuum Pressure < -70kPa |  |  | X | X |  | X | X | X | X |  | X | X
ALIGN | Collided FTM spots are centered at X = 1020 and Y = 540 | X |  | X | X |  |  | X | X |  |  | X | X
ALIGN | FTM adjustment cover turned back, Center Screw Locked |  |  | X | X |  |  | X | X
ALIGN | Side A and SideB Tilt Motor Start Positions Updated and Saved in Config Editor |  |  | X | X |  |  | X | X |  | X | X | X
ALIGN | Coarse XY offset completed | X |  | X | X |  |  | X | X |  | X | X | X
ALIGN | PWM Laser Power Calibration completed |  |  | X | X |  |  | X | X
ALIGN | FineXY Offset completed |  |  | X | X |  | X | X | X | X | X | X | X
ALIGN | X-Tilt Calibration (Ubertarget) completed | ? |  | X | X | [6 empty] | X | X
ALIGN | (4) M5 CAM lockout screws torqued to 80 in-lbs | X |  | X | X |  | X |  | X
ALIGN | (1) M4 CAM lockout screw torqued to 24 in-lbs | X |  | X | X |  | X |  | X
ALIGN | RFID test completed | X |  | X | X | [6 empty] | X | X
ALIGN | X-Tilt Calibration (UBIRD) completed | X |  | X | X |  | X | X | X
ALIGN | BIRD Tilt Motor Start Positions Updated and Saved in Config Editor | X |  | X | X |  |  | X | X |  | X
ALIGN | (3) CAM face mount screws torqued to 50 in-lbs | ? |  |  |  |  | X |  | X |  |  | ? | ?
THERM | Ensure slab is installed and power cable and Ethernet cable is connected | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Ensure USB cable is connected to HCT | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Ensure USB cable is connected to SBC on instrument | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Ensure NVSS initialization completes | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Ensure Instrument Serial number is entered | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Ensure to Home clamps and load HCT | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Ensure vacuum engages at < -75 kPa | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Perform Calibration | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Perform Verification | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Unload HCT from FCH | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Ensure to Home clamps and load HCT | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Ensure vacuum engages at < -75 kPa | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Perform Calibration | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Perform Verification | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Unload HCT from FCH | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Ensure to remove HCT when Calibration and Verification is completed | ? | [9 empty] | ? | ? |  |  |  |  | X | X
THERM | Ensure to Close Door and Exit application | ? | [9 empty] | ? | ? |  |  |  |  | X | X
FIT | Mock Run | ? | [15 empty] | ? | ? | ?
FIT | FIT (Final Instrument Test) Run | X | [15 empty] | X | X | X
Note on B3: Liao, Uland: Initialization encompasses homing of instrument, and chiller ramping which is not necessary in most cases. When not applicable, needs to be noted in Work Instructions. Date: 6/30/2016
Note on B17: Liao, Uland: This step is not provided within TD procedure. However, verification of scripts is the preliminary step to identify out-of-specification settings requiring re-alignment. Date: 6/30/2016
Note on B37: Gates, David: Mock and FIT are setup using multiple steps provided within the TDs. When MOCK or FIT run is performed, all ssteps must be performed; not necessary to have line items for each step.. Date: 23 March 2018
Note on B38: Gates, David: Mock and FIT are setup using multiple steps provided within the TDs. When MOCK or FIT run is performed, all ssteps must be performed; not necessary to have line items for each step.. Date: 23 March 2018